## Slide 1
GC carbon calibration for SWING data
Calibration of intensity for a sample in a glass capillary:
Take the same thickness (in mm) for the capillary filled with solvent
GC calibration: K=0.036
Superimpose GC (cm-1) with the calibration curve by Jan Ilavsky (file: GlassyCarbon_T.dat).
Calibration constant K should be close to 0.036.
= 1.16 mm
In Foxtrot macro, use K=0.036 for the calibration constant (instead of K=1)
Thickness is not taken into account in Foxtrot data reduction

## Slide 2
SWING data April 26: GC calibration
= 1.16 mm
In SASView GUI, go to tools/data operation:
Subtract air signal:
-> Compute and save
- Divide by thickness = 1.16 mm:
-> Compute and save
Multiply by calibration constant K=0.036
-> Compute and save
Verify that superimposes with the calibration curve by Jan Ilavsky (file: GlassyCarbon_T.dat)
GC calibration: K=0.036

## Slide 3
For all SWING data
In Foxtrot, use K=0.036 for the calibration constant (instead of K=1) and run again the macro
Calibration of intensity for a sample in a glass capillary:
** Take the same thickness (in mm) for the capillary filled with solvent